**Appendix No. 1A to the Terms of Reference**

**Requirements and technical parameters**

**for the delivery of ammonia purifier**

 **Part 1**

|  |  |  |
| --- | --- | --- |
| **Parameter name** | **Requirement** | **Column to be filled in by contractor** |
| Type |  |  | Specify |
| Manufacturer  |  |  | Specify |
| Country of origin |  |  | Specify |
| Year of production | 2022 | Required | Confirm |
| Device | Factory new, unused | Required | Confirm |
| The device is used to purify gas – ammonia to a minimum purity of 7N, supplied to the reactor for epitaxy of semiconductor layers based on gallium nitride (GaN) by CVD/MOCVD method. | Required | Confirm |
| 1. The device must purify ammonia to a purity above 7N. The purity of ammonia at the output of the purifier for individual impurities must be:
* H2O < 1 ppb,
* O2 < 1 ppb,
* CO2 < 1 ppb,
* H2 < 1 ppb.
 | Required | Confirm |
| 1. The purifier must purify ammonia at a maximum ammonia flow rate of 50 slm. The average ammonia flow rate is also 30 slm.
 | Required | Confirm |
| 1. The lifetime of the purifier must be a minimum of 1 year, with operation of 6h a day, 5 days a week, ammonia purity at the inlet of 5N, and at the flows shown in section 2.
 | Required | Confirm |
| 1. The required pressure at the outlet of the purifier of 2-3 bar.
 | Required | Confirm |
| 1. The purifier must be equipped with a particle filter (preferably 0.003 μm) with a sintered disc at the outlet.
 | Required | Confirm |
| 1. The purifier must be equipped with a moisture sensor.
 | Required | Confirm |
| 1. The purifier must be equipped with a bypass.
 | Required | Confirm |
| 1. Room-temperature purifier required.
 | Required | Confirm |
| 1. The purifier must be equipped with a VCR 1/4" female connection or corresponding adapter.
 | Required | Confirm |
| 1. Training in operation and maintenance of the device conducted in the Ordering Party's laboratory in or possibly online training.
 | Required | Confirm |
| 1. The supplier will provide training in the operation of the device including construction (basic components) and use and operation of the device.
 | Required | Confirm |
| 1. User's manual, technical documentation and software provided as CD, USB or available for download online
 | Required | Confirm |

 **Appendix No. 1B to the Terms of Reference**

**Requirements and technical parameters**

**for the delivery of argon purifier**

**Part 2**

|  |  |  |
| --- | --- | --- |
| **Parameter name** | **Requirement** | **Column to be filled in by contractor** |
| Type |  |  | Specify |
| Manufacturer  |  |  | Specify |
| Country of origin |  |  | Specify |
| Year of production | 2022 | Required | Confirm |
| Device | Factory new, unused | Required | Confirm |
| The device is used to purify gas – argon to a minimum purity of 7N, supplied to the reactor for epitaxy of semiconductor layers based on silicon carbide (SiC) and gallium nitride (GaN) by CVD/MOCVD method. | Required | Confirm |
| The device must purify argon to a purity above 7N. The purity of argon at the output of the purifier for individual impurities must be: * H2O < 1 ppb,
* O2 < 1 ppb,
* CO < 1 ppb,
* CO2 < 1 ppb,
* H2 < 1 ppb,
* N2 < 1 ppb,
* total hydrocarbons < 1 ppb.
 | Required | Confirm |
| The purifier must purify argon at a maximum argon flow rate of 25 slm. The average argon flow rate is also 25 slm. | Required | Confirm |
| The lifetime of the purifier must be a minimum of 3 years, with operation of 6h a day, 5 days a week, argon purity at the inlet of 5N, and at the flows shown in section 2. | Required | Confirm |
| The required pressure at the outlet of the purifier of 4-10 bar. | Required | Confirm |
| The purifier must be equipped with a particle filter (preferably 0.003 μm). | Required | Confirm |
| The purifier must be mounted in a small cabinet and equipped with a bypass. | Required | Confirm |
| Heated getter purifier with a H2 removal stage required. | Required | Confirm |
| The purifier must be equipped with a VCR 1/4" female connection or corresponding adapter. | Required | Confirm |
| Training in operation and maintenance of the device conducted in the Ordering Party's laboratory or possibly online training. | Required | Confirm |
| The supplier will provide training in the operation of the device including construction (basic components) and use and operation of the device. | Required | Confirm |
| User's manual, technical documentation and software provided as CD, USB or available for download online | Required | Confirm |

**Appendix No. 1C to the Terms of Reference**

**Requirements and technical parameters**

**for the delivery of nitrogen purifier**

**Part 3**

|  |  |  |
| --- | --- | --- |
| **Parameter name** | **Requirement** | **Column to be filled in by contractor** |
| Type |  |  | Specify |
| Manufacturer  |  |  | Specify |
| Country of origin |  |  | Specify |
| Year of production | 2022 | Required | Confirm |
| Device | Factory new, unused | Required | Confirm |
| The device is used to purify gas – nitrogen to a minimum purity of 7N, supplied to the reactor for epitaxy of semiconductor layers based on silicon carbide (SiC) and gallium nitride (GaN) by CVD/MOCVD method. | Required | Confirm |
| The unit must purify nitrogen to a purity above 7N. The purity of nitrogen at the outlet of the purifier for individual impurities must be: * H2O < 1 ppb,
* O2 < 1 ppb,
* CO < 1 ppb,
* CO2 < 1 ppb,
* H2 < 1 ppb.
 | Required | Confirm |
| The purifier must purify nitrogen at a maximum nitrogen flow rate of 125 slm. The average nitrogen flow rate is 100 slm. | Required | Confirm |
| The lifetime of the purifier must be a minimum of 1 year, with operation of 6h a day, 5 days a week, nitrogen purity at the inlet of 5N, and at the flows shown in section 2. | Required | Confirm |
| The required pressure at the outlet of the purifier of 4-10 bar. | Required | Confirm |
| The purifier must be equipped with a particle filter (preferably 0.003 μm). | Required | Confirm |
| The purifier must be equipped with a moisture sensor. | Required | Confirm |
| The purifier must be equipped with a bypass. | Required | Confirm |
| Room-temperature purifier required. | Required | Confirm |
| The purifier must be equipped with an auto-regeneration function. | Required | Confirm |
| The purifier must be equipped with a VCR 1/4" female connection or corresponding adapter. | Required | Confirm |
| Training in operation and maintenance of the device conducted in the Ordering Party's laboratory or possibly online training. | Required | Confirm |
| The supplier will provide training in the operation of the device including: construction (basic components), and use and operation of the device. | Required | Confirm |
| User's manual, technical documentation and software provided as CD, USB or available for download online | Required | Confirm |

**Appendix No. 1D to the Terms of Reference**

**Requirements and technical parameters**

**for the delivery of hydrogen purifier**

**Part 4**

|  |  |  |
| --- | --- | --- |
| **Parameter name** | **Requirement** | **Column to be filled in by contractor** |
| Type |  |  | Specify |
| Manufacturer  |  |  | Specify |
| Country of origin |  |  | Specify |
| Year of production | 2022 | Required | Confirm |
| Device | Factory new, unused | Required | Confirm |
| * 1. The device is used to purify gas - hydrogen to a minimum purity of 7N, supplied to the reactor for epitaxy of semiconductor layers based on silicon carbide (SiC) and gallium nitride (GaN) by CVD/MOCVD method.
 | Required | Confirm |
| * 1. The device must purify hydrogen to a purity above 7N. The purity of hydrogen at the output of the purifier for individual impurities must be:
* H2O < 1 ppb,
* O2 < 1 ppb,
* CO < 1 ppb,
* CO2 < 1 ppb,
* N2 < 1 ppb.

The hydrogen purifier must necessarily remove nitrogen. | Required | Confirm |
| * 1. The purifier must purify hydrogen at a maximum hydrogen flow rate of 350 slm. The average hydrogen flow rate is also 250 slm.
 | Required | Confirm |
| * 1. The lifetime of the purifier must be a minimum of 1 year, with operation of 6h a day, 5 days a week, hydrogen purity at the inlet of 5N and at the flows shown in section 2.
 | Required | Confirm |
| * 1. The purrifier must purify hydrogen in both parts of the Epiluvac ER3 reactor, i.e. SiC and GaN.
 | Required | Confirm |
| * 1. The required pressure at the outlet of the purifier of 4-10 bar.
 | Required | Confirm |
| * 1. The purifier must be equipped with a particle filter (preferably 0.003 μm).
 | Required | Confirm |
| * 1. The purifier must be equipped with a moisture sensor.
 | Required | Confirm |
| * 1. The purifier must be equipped with a hydrogen leakage sensor.
 | Required | Confirm |
| * 1. The purifier must be equipped with an emergency shutdown (EMO) button.
 | Required | Confirm |
| * 1. The purifier must be equipped with an N2 purge system.
 | Required | Confirm |
| * 1. The purifier must be equipped with a bypass.
 | Required | Confirm |
| * 1. Heated getter purifier is required.
 | Required | Confirm |
| * 1. The purifier must be equipped with a PLC with a touch screen HMI for temperature display and control, MODBUS option. Manual operation also possible.
 | Required | Confirm |
| * 1. The purifier must be equipped with a VCR 1/4" female connection or corresponding adapter.
 | Required | Confirm |
| * 1. Training in operation and maintenance of the device conducted in the Ordering Party's laboratory or possibly online training.
 | Required | Confirm |
| * 1. The supplier will provide training in the operation of the device including construction (basic components) and use and operation of the device.
 | Required | Confirm |
| * 1. User's manual, technical documentation and software provided as CD, USB or available for download online
 | Required | Confirm |

**Appendix No. 2 to the Terms of Reference**

**Contractor: Date:** ………………………………………

…………………………………………….

*……………………………………………*

**Statement of the Contractor**

**submitted pursuant to Art. 125, item 1 of the Act of September 11 2019**

**Public Procurement Law (hereinafter: the PPL Act),**

**THE GROUNDS FOR THE EXCLUSION OF ARTICLE 7 item 1 OF THE ACT**

**ON SPECIAL SOLUTIONS IN THE SCOPE OF COUNTERACTING THE SUPPORT OF AGGRESSION AGAINST UKRAINE AND SERVING THE PROTECTION OF NATIONAL SECURITY**

**ON THE CONDITIONS FOR EXCLUSION FROM THE PROCEDURE**

For the purposes of the procedure for award of a public contract, entitled …………………………………………………… *(name of the procedure)*, conducted by Sieć Badawcza Łukasiewicz − Instytut Mikroelektroniki i Fotoniki *(name of the Contracting Authority),* I hereby represent as follows:

**STATEMENTS ON THE CONTRACTOR:**

* I represent that I am not subject to exclusion from the procedure under Art. 108 item 1 and 109 item 1 (4) of the PPL Act and Art. 7 para. 1 (1-3) of the Act of April 13, 2022 on special solutions in the scope of counteracting the support of aggression against Ukraine and serving the protection of national security

or

* that grounds for exclusion from the procedure exist pursuant to Art. …………… of the PPL Act

 *At the same time, I represent that, in connection with these circumstances, I have taken the following corrective actions pursuant to Art. 110 item 2 of the PPL Act: ………………………………………………………………………………….*

**STATEMENT CONCERNING THE ENTITY WHOSE RESOURCES WILL BE REFERRED BY THE CONTRACTOR**

* I certify that the following entity/s, whose resources I am referring to in these proceedings, i.e. …………………………………………………………………….……………………… *(full name of the entity, address, and also depending on the entity: Tax ID/Personal ID, National Court Register/Register of Business Activity Number),* is/are not subject of exclusion from the public Contract awarding procedure.

**STATEMENT CONCERNING A SUBCONTRACTOR WHICH IS NOT AN ENTITY WHOSE RESOURCES WILL BE REFERRED TO BY THE CONTRACTOR**

* I certify that the following entity/s, being the Subcontractor/s: ……………………………………………… *(full name of the entity, address, and also depending on the entity: Tax ID/Personal ID, National Court Register/Register of Business Activity Number)*, is/are not subject of exclusion from the public Contract awarding procedure.

**DECLARATION CONCERNING THE INFORMATION PROVIDED:**

* I declare that all the information provided in the above statements is up-to-date and truthful and has been presented with full awareness of the consequences of misleading the Contracting Authority when providing the information.

I indicate that the documents confirming these facts, referred to in Chapter VII of the ToR, are available in electronic form at the following internet addresses of public and free *(please select):*

* <https://prod.ceidg.gov.pl>
* <https://ems.ms.gov.pl>
* Other register (please provide internet address):[………………………………..](https://ems.ms.gov.pl)

**Appendix No. 3 to the Terms of Reference**

**BID FORM**

for: Łukasiewicz Research Network – Institute of Microelectronics and Photonics

**Part 1**

|  |  |
| --- | --- |
| Subject of the Order | **Delivery of ammonia purifier** |
| 1. | Name and address of the Contractor | ……………………………………………………………………………….………………………………………………………………………………. |
| 2. | Taxpayer’s Identification Number NIPREGON number:or the relevant foreign identification numbers  | ……………………………………………………………………………………………………………………………………………………………..……………………………………………………………………………. |
| 3. | Telephone:E-mail: | ............................................................................................................................................................. |
| 4. | Total net price of the bidVAT rateVAT amountTotal gross price of the bid | **enter:** ......... PLN/EUR/USD/GBP\*say:..................................................................................**enter:** .....................%**enter:** ......... PLN/EUR/USD/GBP\***enter:** ......... PLN/EUR/USD/GBP\*Say: ...................................................... |
| 5. | Time of performing the Contract:Maximum time: 23.12.2022 | **enter:** ........................................  |
| 6. | Warranty period: **at leat 12 months** | **enter:……………………..** *(months)* |
| 7. | Statement | I hereby represent that the selection of our bid **will/will not**\* result in imposing taxation obligation on the Contracting Authority. The following (name or type of the goods or services) will lead to the emergence of a tax obligation**Value without tax:……………………………………………..Name of goods:…………………………………………………..** |
| 8. | Type of Contractor | I declare that I am: Microentrepreneur Small enterprise Medium enterprise A sole proprietorship A natural person who does not conduct business activity Other types: specify ……………………….within the meaning of the Act of March 6, 2018 Entrepreneurs Law.(Note! Mark the "X" in the appropriate box)) |
| 9. | Statement | I hereby represent that I have read the Terms of Reference (together with the Appendices constituting an integral part thereof) and that I accept its content without reservations. |
| 10. | I hereby represent that I have fulfilled the information obligations foreseen in Art. 13 or Art. 14 of the GDPR towards natural persons whose personal data I obtained directly or indirectly in order to apply for awarding a public contract in the present procedure. | **confirm:……………………………………..***In the event if the Contractor does not transfer the personal data other than their own or if the information obligation is excluded pursuant to Art. 13, item 4 or Art. 14, item 5 of the GDPR, the Contractor shall not submit this statement (delete the content of the statement e.g. by crossing out).* |
| 11. | Persons authorised to represent the Contractor | **enter:** ........................................... |
| 12. | Parts of the Contract that will be entrusted to subcontractors | **Specify the parts of the Contract:** .......................**Enter names of companies:** ………………………….. |
| 13. | Documents attached to the bid | 1.2. |

*\* Delete as appropriate*

**Appendix No. 3 to the Terms of Reference**

**BID FORM**

for: Łukasiewicz Research Network – Institute of Microelectronics and Photonics

**Part 2**

|  |  |
| --- | --- |
| Subject of the Order | **Delivery of argon purifier** |
| 1. | Name and address of the Contractor | ……………………………………………………………………………….………………………………………………………………………………. |
| 2. | Taxpayer’s Identification Number NIPREGON number:or the relevant foreign identification numbers  | ……………………………………………………………………………………………………………………………………………………………..……………………………………………………………………………. |
| 3. | Telephone:E-mail: | ............................................................................................................................................................. |
| 4. | Total net price of the bidVAT rateVAT amountTotal gross price of the bid | **enter:** ......... PLN/EUR/USD/GBP\*say:..................................................................................**enter:** .....................%**enter:** ......... PLN/EUR/USD/GBP\***enter:** ......... PLN/EUR/USD/GBP\*Say: ...................................................... |
| 5. | Time of performing the Contract:Maximum time: 23.12.2022 | **enter:** ........................................  |
| 6. | Warranty period: **at leat 12 months** | **enter:……………………..** *(months)* |
| 7. | Statement | I hereby represent that the selection of our bid **will/will not**\* result in imposing taxation obligation on the Contracting Authority. The following (name or type of the goods or services) will lead to the emergence of a tax obligation**Value without tax:……………………………………………..Name of goods:…………………………………………………..** |
| 8. | Type of Contractor | I declare that I am: Microentrepreneur Small enterprise Medium enterprise A sole proprietorship A natural person who does not conduct business activity Other types: specify ……………………….within the meaning of the Act of March 6, 2018 Entrepreneurs Law.(Note! Mark the "X" in the appropriate box)) |
| 9. | Statement | I hereby represent that I have read the Terms of Reference (together with the Appendices constituting an integral part thereof) and that I accept its content without reservations. |
| 10. | I hereby represent that I have fulfilled the information obligations foreseen in Art. 13 or Art. 14 of the GDPR towards natural persons whose personal data I obtained directly or indirectly in order to apply for awarding a public contract in the present procedure. | **confirm:……………………………………..***In the event if the Contractor does not transfer the personal data other than their own or if the information obligation is excluded pursuant to Art. 13, item 4 or Art. 14, item 5 of the GDPR, the Contractor shall not submit this statement (delete the content of the statement e.g. by crossing out).* |
| 11. | Persons authorised to represent the Contractor | **enter:** ........................................... |
| 12. | Parts of the Contract that will be entrusted to subcontractors | **Specify the parts of the Contract:** .......................**Enter names of companies:** ………………………….. |
| 13. | Documents attached to the bid | 1.2. |

*\* Delete as appropriate*

**Appendix No. 3 to the Terms of Reference**

**BID FORM**

for: Łukasiewicz Research Network – Institute of Microelectronics and Photonics

Part 3

|  |  |
| --- | --- |
| Subject of the Order | **Delivery of of nitrogen purifier** |
| 1. | Name and address of the Contractor | ……………………………………………………………………………….………………………………………………………………………………. |
| 2. | Taxpayer’s Identification Number NIPREGON number:or the relevant foreign identification numbers  | ……………………………………………………………………………………………………………………………………………………………..……………………………………………………………………………. |
| 3. | Telephone:E-mail: | ............................................................................................................................................................. |
| 4. | Total net price of the bidVAT rateVAT amountTotal gross price of the bid | **enter:** ......... PLN/EUR/USD/GBP\*say:..................................................................................**enter:** .....................%**enter:** ......... PLN/EUR/USD/GBP\***enter:** ......... PLN/EUR/USD/GBP\*Say: ...................................................... |
| 5. | Time of performing the Contract:Maximum time: 23.12.2022 | **enter:** ........................................  |
| 6. | Warranty period: **at leat 12 months** | **enter:……………………..** *(months)* |
| 7. | Statement | I hereby represent that the selection of our bid **will/will not**\* result in imposing taxation obligation on the Contracting Authority. The following (name or type of the goods or services) will lead to the emergence of a tax obligation**Value without tax:……………………………………………..Name of goods:…………………………………………………..** |
| 8. | Type of Contractor | I declare that I am: Microentrepreneur Small enterprise Medium enterprise A sole proprietorship A natural person who does not conduct business activity Other types: specify ……………………….within the meaning of the Act of March 6, 2018 Entrepreneurs Law.(Note! Mark the "X" in the appropriate box)) |
| 9. | Statement | I hereby represent that I have read the Terms of Reference (together with the Appendices constituting an integral part thereof) and that I accept its content without reservations. |
| 10. | I hereby represent that I have fulfilled the information obligations foreseen in Art. 13 or Art. 14 of the GDPR towards natural persons whose personal data I obtained directly or indirectly in order to apply for awarding a public contract in the present procedure. | **confirm:……………………………………..***In the event if the Contractor does not transfer the personal data other than their own or if the information obligation is excluded pursuant to Art. 13, item 4 or Art. 14, item 5 of the GDPR, the Contractor shall not submit this statement (delete the content of the statement e.g. by crossing out).* |
| 11. | Persons authorised to represent the Contractor | **enter:** ........................................... |
| 12. | Parts of the Contract that will be entrusted to subcontractors | **Specify the parts of the Contract:** .......................**Enter names of companies:** ………………………….. |
| 13. | Documents attached to the bid | 1.2. |

*\* Delete as appropriate*

**Appendix No. 3 to the Terms of Reference**

**BID FORM**

for: Łukasiewicz Research Network – Institute of Microelectronics and Photonics

Part 4

|  |  |
| --- | --- |
| Subject of the Order | **Delivery of hydrogen purifier** |
| 1. | Name and address of the Contractor | ……………………………………………………………………………….………………………………………………………………………………. |
| 2. | Taxpayer’s Identification Number NIPREGON number:or the relevant foreign identification numbers  | ……………………………………………………………………………………………………………………………………………………………..……………………………………………………………………………. |
| 3. | Telephone:E-mail: | ............................................................................................................................................................. |
| 4. | Total net price of the bidVAT rateVAT amountTotal gross price of the bid | **enter:** ......... PLN/EUR/USD/GBP\*say:..................................................................................**enter:** .....................%**enter:** ......... PLN/EUR/USD/GBP\***enter:** ......... PLN/EUR/USD/GBP\*Say: ...................................................... |
| 5. | Time of performing the Contract:Maximum time: 23.12.2022 | **enter:** ........................................  |
| 6. | Warranty period: **at leat 12 months** | **enter:……………………..** *(months)* |
| 7. | Statement | I hereby represent that the selection of our bid **will/will not**\* result in imposing taxation obligation on the Contracting Authority. The following (name or type of the goods or services) will lead to the emergence of a tax obligation**Value without tax:……………………………………………..Name of goods:…………………………………………………..** |
| 8. | Type of Contractor | I declare that I am: Microentrepreneur Small enterprise Medium enterprise A sole proprietorship A natural person who does not conduct business activity Other types: specify ……………………….within the meaning of the Act of March 6, 2018 Entrepreneurs Law.(Note! Mark the "X" in the appropriate box)) |
| 9. | Statement | I hereby represent that I have read the Terms of Reference (together with the Appendices constituting an integral part thereof) and that I accept its content without reservations. |
| 10. | I hereby represent that I have fulfilled the information obligations foreseen in Art. 13 or Art. 14 of the GDPR towards natural persons whose personal data I obtained directly or indirectly in order to apply for awarding a public contract in the present procedure. | **confirm:……………………………………..***In the event if the Contractor does not transfer the personal data other than their own or if the information obligation is excluded pursuant to Art. 13, item 4 or Art. 14, item 5 of the GDPR, the Contractor shall not submit this statement (delete the content of the statement e.g. by crossing out).* |
| 11. | Persons authorised to represent the Contractor | **enter:** ........................................... |
| 12. | Parts of the Contract that will be entrusted to subcontractors | **Specify the parts of the Contract:** .......................**Enter names of companies:** ………………………….. |
| 13. | Documents attached to the bid | 1.2. |

*\* Delete as appropriate*

**Appendix No. 4 to the Terms of Reference**

**OBLIGATION OF AN ENTITY TO PROVIDE THE CONTRACTOR WITH THE RESOURCES NECESSARY FOR PERFORMING THE CONTRACT**

Having read the content of the announcement of the Contract and the Terms of Reference for the procedure for awarding a public Contract conducted in form of an open tender by ……………………………………………………….. for: **………………………….**, we:……………… (*first name and surname of the person signing the statement)* hereby represent that we are duly authorised to perform this action, acting on behalf of:…………………………………………………………. (*enter the name of the entity providing resources)* with its seat in: ………………………. ……………………. *(enter the address of the entity providing resources)* we hereby commit ourselves to provide access to the resources in our possession necessary to realise the Contract, to: ………………. *(enter the name)* with its seat in: ……………, hereinafter referred to as the Contractor.

1. The scope of the resources available to the Contractor:
	1. ………………………………………………………………………………………………

*(please specify the provided resources)*

* 1. ………………………………………………………………………………………………

*(please specify the provided resources)*

1. Manner of using the resources in the realisation of the Contract: …………………………………………………………………………………………………................................................
2. Scope and duration of our participation in performing the Contract: ………………………………………………………………………………………………….................................................
3. We will provide the following services being parts of the subject of the Contract:

……………………………………………………………………………………………………….................................................

Due to the above, we hereby provide these resources to the disposal of the Contractor who may use them in the realisation of the subject of the Contract (if its bid is selected and the Contract granted in the said procedure).

**Appendix No. 5 to the Terms of Reference**

**Contractors who jointly apply for a contract:**

………………………………………

………………………………………

 (Full name/Company name, address)

**STATEMENT OF THE CONTRACTOR ABOUT THE SCOPE OF PERFORMANCE OF THE CONTRACT BY THE CONTRACTORS WHO JOINTLY APPLY FOR A CONTRACT**

submitted on the basis of art. 117 item 4 of the Act of September 11, 2019 - Public Procurement Law (Official Journal of 2019, item 2019 with later changes)

in the procedure for awarding a public Contract for

**………………………………………………………………………………………..**

we declare that:

1. Contractor …………………………………………………………………………………………………….…………….…..

*(name and address name and address of one of the Contractors who jointly apply for a contract, e.g. a consortium member or partner in a civil partnership)*

will perform the following supplies, services or construction works \*:…………………………………………

1. Contractor ……………………………………………………………………………………………….……………..……..

*(name and address name and address of one of the Contractors who jointly apply for a contract, e.g. a consortium member or partner in a civil partnership)*

will perform the following supplies, services or construction works \*:…………………………………………

NOTE!

It should be indicated which supplies/services/construction works \* included in scope of the subject of the contract are performed by individual contractors (consortium members or partners of a civil partnership)

**STATEMENT REGARDING THE INFORMATION PROVIDED**

I/We\* hereby represent that any and all information provided in this Statement are up-to-date and accurate and that they I/we are fully aware of the consequences of providing the Contracting Authority with misleading information.

*\* Delete as appropriate*

**Appendix No. 7 to the Terms of Reference**

**STATEMENT**

STATEMENT OF THE CONTRACTOR CONFIRMING THE VALID OF INFORMATION PROVIDED IN THE STATMENT SPECIFIED IN ART. 125 OF THE PPL ACT, ON THE GROUNDS FOR EXCLUSION

Contractor :……………………………………………………………………………………………

…………………………………………………………………………………………………….……

For the purposes of procurement procedure, the subject of which is
……………………………………………………………………………………………………………………………………………………………..,

We declar that the information provided in the statement placed pursuant to Art. 125, item 1 of the PPL Act, with respect to the ground for exclusion from procedinds, stipalated in:

1. **art. 108 ust. 1** of the PPL Act
2. **art. 7 para. 1 (1-3)**  of the Act of April 13, 2022 on special solutions in the scope of counteracting the support of aggression against Ukraine and serving the protection of national security

**are up-tp-date / are out-of-date\***

NOTE: In the previously provided information is not up-to-date, please submit additional information in this regard, in particular specify which data the change relatesto and indicate its scope

*\*delete as appropriate*